## Slide 2: Method Parameters

## Slide 3
Data acquisition parameters
Ion Mode: positive, negative
Instrument Mode: linear, reflector
Instrument range: mass range
Low mass gate: on/off, cutoff mass
Total scans: no. of laser shots averaged
Accelerating voltage: 20-25 kV
Delay time: time between laser flash and ion extraction Grid voltage: expressed as a % of Accel. Voltage
Guide wire voltage: ditto
Parameters that require optimization in linear/reflector:

## Slide 4
Optimization Strategy

## Slide 5
Laser Power
Affects S/N and Resolution
A different power setting will be needed for 3 vs 20 Hz acquisition rate
A different setting is needed for different matrices and sample types
Excessive laser power will result in saturated peaks with poor resolution and high sample consumption

## Slide 6
Bin Size
(Data Collection Interval)
Data collected at 1 ns intervals
Baseline resolution between adjacent peaks.
Incomplete resolution between same peaks
Data collected at 4 ns intervals
Determines the time interval between subsequent acquired data points. Increasing the number of data points by sampling more frequently can increase resolution for a given mass range but also increases the size of the data file.

## Slide 7
Laser
Sample
plate
Beam guide wire
Reflector
Linear
detector
Variable
voltage
grid
Reflector
detector
Laser
Attenuator
Ground grid
Guide Wire and Accelerating voltages

## Slide 8
Accelerating Voltage
Accelerating Voltage
Linear
Reflector
Increasing can improve sensitivity for higher mass compounds (>25-20 kDa). Typical 20 or 25 kV.
Decreasing can improve resolution on compounds <2 kDa.
Increasing can improve sensitivity
The accelerating voltage determines the kinetic energy of the ions when they reach the detector. Efficiency of detection increases somewhat with higher ion energy.
A lower accelerating voltage provides more data points across a peak (ions move slower) for better resolution.
Max: 25 kV

## Slide 9
Guide wire voltage
Guide wire voltage %
Decreasing can improve resolution
Increasing can improve sensitivity for higher masses
To obtain maximum resolution in reflector mode, set the guide wire to 0%. This can then be adjusted up to 0.02%.
In linear mode <2 kDa, a setting of 0.05-0.1% is adequate.
In linear mode >20 kDa, start with 0.3% and decrease as needed.
Note: New DE STR instruments have a lens in place of the guide wire, so no adjustment is necessary.

## Slide 10
Delayed Extraction
Ref: W.C. Wiley and I.H. McLaren, Rev. Sci. Instrum. (1953) 26, 1150-1157.
When ions are formed in MALDI they have a range of translational kinetic energies due to the ionization process. This leads to peak broadening. By forming ions in a weak electric field, then applying a high voltage extracting field only after a time delay, the effect of this energy spread can be minimized when used in conjunction with an appropriate potential gradient.
Field gradients are formed and controlled in the ionization region by the voltages applied to the sample plate and the variable voltage grid.

## Slide 11
+20 kV
+18 kV (90%)
Potential
Gradient
Ionization Region
Sample plate at accelerating voltage
Variable voltage at a % of the accelerating voltage
Ground grid
The variable voltage works together with the accelerating voltage to create a potential gradient in the ionization region near the target. It and the delay time must be adjusted to obtain optimum resolution for a given mass range.
to Flight Tube
Extraction Voltages
Notes: The 4700 software automatically adjusts the gradient and delay time based on the desired Focus Mass

## Slide 12: Pulse Delay Time with Delayed Extraction Technology
time
Extraction delay time (25-1000 ns)
Laser pulse
Accelerating Voltage
kV
Variable Voltage

## Slide 13
The problem: Peaks are broad in MALDI-TOF spectra with continuous extraction (=poor resolution).
The cause: Ions of the same mass coming from the target have different Kinetic Energy (velocity) due to the ionization process.
+
+
+
Sample+matrix on target
Ions of same mass but different velocities (KE)
Ion Extraction

## Slide 14
+
+
+
The result: Ions of the same mass extracted immediately out of the source with a uniform accelerating voltage will have a broad spread of arrival times at the detector resulting in a broad peak with poor resolution.
Ion Extraction
Detector

## Slide 15
The solution: Delayed Extraction (DE)
Ions are allowed to spread out away from the plate during an appropriate time delay prior to applying the accelerating voltage
Delayed Extraction (DE)
+
+
+
The position of an ion in the source after the pulse delay will be correlated with its initial velocity or kinetic energy
Ions of same mass but different velocities (KE)

## Slide 16
Velocity Focusing with DE
+
+
+
Ions of same mass, different velocities
1: No electric field. Ions spread out during delay time.
2: Field applied. Gradient accelerates slow ions more than fast ones.
+
+
+
3: Slow ions catch up with faster ones at the detector.
0 V
+20 kV
0 V
+
+
+
Detector
+18 kV
Sample Plate
Variable Voltage Grid
Notes: This is known as Velocity Focusing

## Slide 17
m/z
6130
6140
6150
6160
6170
10600
10800
11000
11200
11400
11600
m/z
Linear mode
Reflector mode
continuous
extraction
R=125
delayed
extraction
R=1,100
delayed
extraction
R=11,000
continuous
extraction
R=650
Delayed Extraction:
Resolution Improvements circa 1996
Sample: mixed base DNA 36-mer
Sample: mixed base DNA 20-mer

## Slide 18: Optimizing grid voltage % and delay time
Grid Voltage % and Delay time are interactive parameters.
For each grid voltage % there is an optimal delay time.
Increments of 0.3% in grid % or 50 ns in delay may give significantly different performance.
The general trends are shown in the table above.

## Slide 19
0
200
400
600
87
88
89
90
91
92
93
94
95
96
Grid Voltage (%)
Pulse Delay (ns)
1000
2000
5000
15000
25000
m/z=50000
Typical curves of optimum delay time as a function of grid voltage in linear mode

## Slide 20
Voyager DE/RP/PRO/STR Delay Time and Grid Voltage %

## Slide 21
Optimizing a Delayed Extraction Method
1. Start with a standard method on a known sample.
2. Find an adequate laser setting that gives good peak intensity without saturation.
3. Set the guide wire voltage for best sensitivity (peak intensity and/or S/N). Use lowest practical guide setting.
4. Optimize the grid voltage or the delay time, leaving the other unchanged. These parameters are interactive, so each must be optimized separately. Optimize for highest resolution.
5. Recheck 3-4, see if you get same results.

## Slide 23
Calibration
Voyager Training Class

## Slide 24: Calibration Equations
T = to + A m/z + ( higher order terms)
Where
to = difference in time between the start of analysis and the time of ion extraction.
A = effective length (mm) mo
Where mo = 1 dalton mass in SI units
e = charge of electron in SI units
Effective length = length of flight tube corrected for ion acceleration
e
x
X 10 9
 Accelerating Voltage (kV)

## Slide 25: Initial Velocity Correction
Initial velocity is the average speed at which matrix ions desorb.
The initial velocity (m/s) has been calculated for different matrices. The calibration equation can be corrected for matrix initial velocity (one of the higher order terms).
Externally calibrated samples must be in the same matrix as their calibrant.
CHCA 300 m/s
Sinapinic acid 350 m/s
DHB 500 m/s
3-HPA 550 m/s
Ref:Juhasz,P.,M.Vestal, and S.A.Martin. J.Am.Soc.Mass Spectrom.,1997,8,209-217

## Slide 26: Calibration Equations
A default calibration uses a multiparameter equation that estimates values for tº and A from instrument dimensions.
Default calibration is applied to the mass scale if no other calibration is specified.

## Slide 27: Calibration Equations
Internal calibration uses a multiparameter equation that calculates values for tº and A using the known mass of the standard(s). This corrects the mass scale.
A multi-point calibration calculates tº and A by doing a least-squares fit to all of the standards.
A two point calibration calculates tº and A from the standards. A one point calibration calculates A from the standard and uses tº from the default calibration.

## Slide 28: Internal Calibration
A one-, two- or multi-point calibration using known peak masses that are within the spectrum to be calibrated.
The standards should bracket the mass range of interest. The signal intensities of the standards should be similar to those of the samples.
The calibration equation is saved within the data file and can be exported as a *.cal file to the acquisition method or to another data file.

## Slide 29
Useful Calibration Standards
Sequazyme Mass Standards Kit: P2-3143-00
Sequazyme BSA Test Standard: 2-2158-00
Voyager IgG1 Mass Standard: GEN 602151
Other useful high mass calibrants:
Cytochrome C: 12,231
Bovine Trypsin: 23,291
Carbonic Anhydrase: 29,024
Bakers Yeast Enolase: 46,672

## Slide 30
756.4732
807.438
3
893.5128
944.4889
1068.6892
1412.8272
1471.7961
1627.9507
1789.8437
1821.9344
1840.928
1
1875.9875
1948.0391
2039.1459
2124.0419
2349.1390
2441.1330
2471.2187
2583.3076
2973.4467
3178.6034
3187.7327
3490.7970
0
5000
10000
15000
1000
2000
3000
MALDI TOF mass spectrum of the
tryptic
digest of yeast
enolase
(in
a
-
cyano
-4-
hydroxy cinnamic
acid matrix) acquired in reflector mode. Peaks at m/z 756.47 and
3187.73 were used as internal
calibrants
.
Two point Internal Calibration

## Slide 31
Enolase
AVSKVYARSVYDSRGNPTVEVELTTEKGVFRSIVPSGASTGVHEALEMR
D
GDKSKWMGKGVLHAVKNVNDVIAPAFVK
ANIDVKDQKAVDDFLISLDGTANKSKLGANAILGVSLAASR
AAAAEKNVPLYKHLADL
SKSKTSPYVLPVPFLNVLNGGS
HAGGALALQEFMIAPTGAKTFAEALRIGSEVYHNLKSLTKKRYGASAGNVGDEGGVAPNIQTAEEALDLIVDAIKAAGHD
GKV
KIGLDCASSEFFKDGKYDLDFKNPNSDKSKWLTGPQLADLYHSLMKRYPIVSIEDPFAEDDWEAWSHFFKTAGIQIV
ADDLTVTNPKRIATAIEK
KAADALLLKVNQIGTLSESIKAAQDSFAAGWGVMVSHRSGETEDTFIADLVVGLRTGQIKTG
APARSERLAKLNQLLRIEEELGDNAVFAGENFHHGDKL
m/z
MH+
Delta(
ppm
)
Start/end
Peptide Sequence
756.4732
756.4732
0.0028
415-420
(K)LNQLLR(I)
807.4382
807.4365
2.1327
180-187
(K)TFAEALR(I)
893.5131
893.5209
-8.7031
1-8
(-)AVSKVYAR(S)
944.4884
944.4914
-3.1415
403-411
(K)TGAPARSER(L)
1068.6889
1068.6893
-0.4105
412-420
(R)LAKLNQLLR(I)
1412.8272
1412.8225
3.2999
106-120
(K)LGANAILGVSLAASR(A)
1471.8047
1471.7981
4.4748
398-411
(R)TGQIKTGAPARSER(L)
1627.9507
1627.9495
0.7191
104-120
(K)SKLGANAILGVSLAASR(A)
1789.8439
1789.8444
-0.2862
363-380
(K)AAQDSFAAGWGVMVSHR(S)
1821.9345
1821.9234
6.0739
381-397
(R)SGETEDTFIADLVVGLR(T)
1840.9284
1840.9227
3.0842
32-49
(R)SIVPSGASTGVHEALEMR(D)
1875.9879
1875.9816
3.3490
15-31
(R)GNPTVEVELTTEKGVFR(S)
1948.0390
1948.0292
5.0121
180-197
(K)TFAEALRIGSEVYHNLK(S)
2039.1460
2039.1290
8.3612
121-140
(R)AAAAEKNVPLYKHLADLSK(S)
2124.0417
2124.0461
-2.0566
9-27
(R)SVYDSRGNPTVEVELTTEK(G)
2441.1344
2441.1373
-1.2055
421-444
(R)IEEELGDNAVFAGENFHHGDKL(-)
2471.1987
2471.2200
-8.6300
32-55
(R)SIVPSGASTGVHEALEMRDGDKSK(W)
2583.3073
2583.3055
0.7080
9-31
(R)SVYDSRGNPTVEVELTTEKGVFR(S)
2973.4477
2973.4563
-2.8758
32-59
(R)SIVPSGASTGVHEALEMRDGDKSKWMGK(G)
3178.6048
3178.5922
3.9794
415-444
(K)LNQLLRIEEELGDNAVFAGENFHHGDKL(-)
3187.7327
3187.7327
0.0103
89-120
(K)AVDDFLISLDGTANKSKLGANAILGVSLAASR(A)
3490.8160
3190.8083
2.2081
412-444
(R)LAKLNQLLRIEEELGDNAVFAGENFHHGDKL(-)
Fig. 2
Summary of
enolase
peptides identified by MALDI TOF. Upper : expected sequence with confirmed
sequences underlined. Lower : detailed mass data for matched peptides.
High Mass Accuracy Achieved with a Two Point Internal Calibration

## Slide 32: External Calibration
Calibration from one standard applied to another nearby sample. The closer the standard is to the sample spot, the better the calibration, but not as good as internal calibration.
Central External Standard
Sample wells
Close External Standard

## Slide 33: Using an external calibration file in the ICP
Specify the External Calibration file here
If you specify an external calibration file in the ICP, all data files will have that calibration applied automatically as they are acquired

## Slide 34
Voyager
Instrument Control Panel
Voyager Training Class

## Slide 35: Elements of the Control Panel
Data storage
page
Sample plate
page
Toolbar
Status bar
Spectrum view
Instrument
settings page
Instrument
status page
Output
window
Calibration file
Instrument Mode
Laser Step Control
Control Mode

## Slide 36: System Status Display
Green = OK/ON
Yellow = Fault
Gray = Off
Status Window
Status Bar

## Slide 37
Selecting the Sample Plate Type

## Slide 38
Select the plate ID, or input a new name
Select the *.plt file
Date when (if) the plate was last aligned
Selecting the sample plate type
Date when (if) a plate optimization file was created

## Slide 39
Sample Plate Alignment

## Slide 40: Simple Acquisition
Step 1: Open acquisition method
Step 2: Specify data file name
Step 3: Move to sample position
Step 4: Acquire/view data

## Slide 41: Opening Instrument Setting (Method Files)*.BIC

## Slide 42: Standard Linear Methods
Angiotensin_linear.bic 500-2500
ACTH_linear.bic 500-5,000
Insulin_linear.bic 1000-10,000
Myoglobin_linear.bic 1,000-25,000
BSA_linear.bic 2,000-100,000
IgG_linear.bic 10,000-200,000

## Slide 43: Standard Reflector Methods
Angiotensin_reflector.bic 500-2,500
ACTH _reflector.bic 1,000-4,000
Insulin _reflector.bic 2,500-7,000
Thioredoxin _reflector.bic 1,000-15,000
psd_precursor.bic variable
Angiotensin_psd.bic 1,296.7
Angiotensin_auto psd.bic 1,296.7

## Slide 44
Data storage
Enter a Root filename
Enter a sample description
Specify data directory here or create a new directory under the File menu

## Slide 45
To save Data

## Slide 46
To go to Data Explorer and open the Saved Data File

## Slide 47: Sample position control
Laser power
setting
Right mouse click toggles between normal and expanded views

## Slide 48
Sample View (con’t)
Laser Intensity Controls
Point and Click with the mouse to move to new x,y coordinates or use Joystick
Coarse laser control
Slider laser control
Fine laser control
Fine and Coarse step sizes are set up in Hardware Config. / Laser
Instrument

## Slide 49: Spectrum Accumulation in Manual Mode
Step 1: Acquire “a few” laser shots
Step 2: Inspect Current Spectrum
Step 3: Optionally use the calculators
Step 4: If spectrum is satisfactory
add to the accumulation
buffer.
Q: When to use?
A: When the user cannot afford “bad” scans

## Slide 50
Resolution and Signal to Noise Calculators

## Slide 51: Resolution and S/N calculators during acquisition
2466.2 (R8566, S146.3)
Tools
Tools

## Slide 52: Instrument Settings / Mode
Instrument
mode tab
Digitizer
tabs
TIS / Reflector
Tune Ratio
Instrument Modes
Voltage settings
Acquisition control mode
Mass range settings
Calibration mode

## Slide 53
Mode / Digitizer Setup

## Slide 54
Advanced Tab from Instrument Setting

## Slide 55
Standard Voyager Acquisition Methods
The following pages contain details of the standard methods (*.bic files) for Voyager DE, DE-PRO and DE-STR. These files are usually found in the C drive of the Voyager computer in Voyager/Data/Installation. The instrument settings shown in these tables are only starting points and may be different than the actual settings required to achieve a given specification. The Grid
Voltage % and Guide Wire % settings are the most critical for method optimization.
The settings required to optimize any method varies from one instrument to the next, thus a .bic file copied from another instrument will not necessarily work well on yours without additional fine-tuning.
Keep at least one copy of your optimized .bic files in a write-protected folder. Create a working copy of these files for daily use.