## Example
Project Name: |  | Duke Buck Station
Project Location: |  | Dukeville Road, Salisbury
Project Number: |  | 7126-16-032A
Water Quality Meter Calibration
Date: | 2016-09-28
Personnel: | David Klemm, Brant Alyea
Equipment ID: | YSI Flow-through cell with meter
Model: | YSI Professional Series
Serial #: | 13K100921
| Time of Calibration | Battery Check (Pass/Fail) | Pre/Post Calibration or Midday Check | pH Buffer 4 (S.U.) | pH Buffer 7 (S.U.) | pH Buffer 10 (S.U.) | DO (mg/L) | Specific Conductivity (µmhos/cm) | ORP (mV)
| 7:50 | Pass | Pre Cal | 3.97 | 7.01 | 10 | 0.13 | 997 | 252.4
Midday
| 15:30 | Pass | Post Cal | 4.13 | 6.94 | 9.97 | 0 | 997 | 249.9
pH Buffer 4 |  |  | DO Solution (0 mg/L) |  |  |  | Comments:
Lot #: | 6GA755 |  | Lot #: | 201046-2
Expiration date: | 2018-01-01 |  |  | Mix date: | 2016-09-26
pH Buffer 7 |  |  | ORP Solution (257.6 mV) |  | 257.6
Lot #: | 6GD380 |  | Lot #: | NA
Expiration date: | 2018-04-01 |  |  | Mix date: | 2016-09-19
pH Buffer 10 |  |  | Specific Conductvity Solution (1000 µmhos/cm)
Lot #: | 6GA660 |  | Lot #: | 6GF129
Expiration date: | 2018-04-01 |  |  | Expiration date: | 2017-06-01
Turbidity Meter Calibration
| Date: | 2016-09-28
| Personnel: | David Klemm, Brant Alyea
| Equipment ID: | Micro TPW Turbidimeter
| Model: | Micro TPW Turbidimeter
| Serial #: | 201505263
|  | Time of Calibration | Battery Check (Pass/Fail) | Pre/Post Calibration or Midday Check | Turbidity Solution
[5 empty] | 0.02 NTU | 10 NTU | 1000 NTU
|  | 7:50 | Pass | Pre Cal | 0.02 | 10.01 | 1000
Midday
|  | 15:12 | Pass | Post Cal | 0.01 | 9.99 | 1003
| TurbidIty - 10 NTU Solution |  |  |  | Comments:
| Lot #: | 50539
| Expiration date: | 2017-05-01
TurbidIty - 0.02 NTU Solution
| Lot #: | 50501
| Expiration date: | 2017-05-01
TurbidIty - 1000 NTU Solution
| Lot #: | 50503
| Expiration date: | 2017-05-01